July 18, 1989

Program LABCAL

Programmer: R. Millard

Computer: VAX/PC

Language: fortran

LINK LABCAL

Purpose: The LABCAL program evaluates the polynomial coefficients

relating the CTD pressure, temperature, and conductivity sensor

observations to laboratory standard values.  Polynomials up to third order

can be evaluated for each sensor, although in practice for the E G & G

Mark III CTD, only pressure is found to require a 3rd while temperature

normally requires a 2nd order and conductivity only a first order fit.

Input:

The calibration data has the following free field input form:

temperature/conductivity

sequence #, T_stnd, C_stnd, T-instr, C-instr/

.

.

.

.

any negative temperature or conductivity value less than -3. is

removed from the fitting.  Separate counters are used for temperature and

conductivity.

pressure

-99,room temperature C/  -99 = start pressure, pres. cal temperature

Plab, P-instr

.

.

.

.

The laboratory calibrations values of temperature/conductivity and

pressure can be merged together provided pressure is last.   The -99

start pressure flag distinquishes these data types.

The laboratory standard pressure is recorded in psi (pounds per square

inch) while the CTD calibration is required in decibars ( 10^4 pascals).

The conversion of psi to decibars requires multiplying the pressure

standard by .6894734.  Other corrections to the pressure measurement

include local gravity, fluid head, temperature and pressure distortions to

the dead weight tester piston (see WHOI TR 74-89).

The CTD conductivity measurement is corrected for conductivity cell

temperature distortion using ` = -6.5 E-5 for alumina before fitting.  The

CTD temperature observation is used.

Reference:

Fofonoff,  N.  P.,  S.  P.  Hayes,  and R.  C. Millard   1974

WHOI/Brown  CTD  Microprofiler:  Methods of calibration and  data

handling.  Woods Hole.Oceanographic Institution Tech. Rep. 74-89.

Program operation:

The following are program prompts and response.

The program requires the operator know the input calibration file name.

This file must have an extension of *.CAL.  The fluid head pressure bias

"FHBIAS"  in decibars of the pressure set up is defaulted to 0.0 dbars and

the "Pslope" pressure slope of the deck unit display is defaulted to .1

dbars.

R LABCAL

PROGRAM VERSION    LABCAL  891208

PROGRAM LABCAL Vers. 1

ENTER NAME OF *.CAL  FILE TO FIT

9c88ncv

LABFIT: FHBIAS,Pslope

.4,.1/

# T, C, P PTS=            7           7          36          12

ENTER VAR,ORDER FIT=-1 for DELTA file

1=T, 2=C; 3=Pdwn;4=Pup; 5= Pall; 6=New File; 7=exit

1,-1/  TEMPERATURE DIFFERENCE FILE

ENTER VAR,ORDER FIT=-1 for DELTA file

1=T, 2=C; 3=Pdwn;4=Pup; 5= Pall; 6=New File; 7=exit

1,2/ QUADRATIC TEMPERATURE FIT

ENTER VAR,ORDER FIT=-1 for DELTA file

1=T, 2=C; 3=Pdwn;4=Pup; 5= Pall; 6=New File; 7=exit

2,1/  LINEAR CONDUCTIVITY FIT

ENTER VAR,ORDER FIT=-1 for DELTA file

1=T, 2=C; 3=Pdwn;4=Pup; 5= Pall; 6=New File; 7=exit

3,3/ DOWN CUBIC PRESS.

ENTER VAR,ORDER FIT=-1 for DELTA file

1=T, 2=C; 3=Pdwn;4=Pup; 5= Pall; 6=New File; 7=exit

4,3/ UP CUBIC PRESS.

ENTER VAR,ORDER FIT=-1 for DELTA file

1=T, 2=C; 3=Pdwn;4=Pup; 5= Pall; 6=New File; 7=exit

3,-1/ DOWN PRESS. DIFFERENCE

ENTER VAR,ORDER FIT=-1 for DELTA file

1=T, 2=C; 3=Pdwn;4=Pup; 5= Pall; 6=New File; 7=exit

7,0/  EXIT PROGRAM

Sample fit parameter file

$ type 9c88nv.fit

LABFIT: FHBIAS,Pslope

0.4000000      0.1000000

MEAN CCR= 0.999794E-03   PRS ROOM TEMP=  22.00

PROGRAM VERSION    LABCAL  891208

DISK FILE = 9c88nv.cal

VARIABLE=TEMPERATURE

7 DATA POINTS          ORDER OF POLY  2

POLY COEF =0.953261E-02 0.499906E-03 0.104558E-11

INSTRUMENT      OBSERVED      CALCULATED      DIFFERENCE

1         48002.0          24.008200      24.008415      -0.000216

2         40006.0          20.011000      20.010435       0.000565

3         29998.0          15.006200      15.006646      -0.000446

4         20004.0          10.010000      10.010065      -0.000065

5         10034.0           5.025900       5.025692       0.000208

6         6002.00           3.010100       3.010005       0.000095

7         998.000           0.508300       0.508440      -0.000140

MEAN DEVIATION =   0.289508E-06

STANDARD DEVIATION =   0.326754E-03

PROGRAM VERSION    LABCAL  891208

DISK FILE = 9c88nv.cal

VARIABLE=CONDUCTIVITY

7 DATA POINTS          ORDER OF POLY  1

POLY COEF =0.396781E-02 0.999686E-03

INSTRUMENT      OBSERVED      CALCULATED      DIFFERENCE

1         50623.7          50.612202      50.611794       0.000408

2         46625.7          46.615200      46.614994       0.000206

3         41754.8          41.745300      41.745644      -0.000343

4         37058.6          37.050301      37.050915      -0.000614

5         32560.1          32.553200      32.553795      -0.000595

6         30798.6          30.792601      30.792854      -0.000254

7         28661.1          28.657301      28.656096       0.001205

MEAN DEVIATION =   0.190735E-05

STANDARD DEVIATION =   0.655196E-03

PROGRAM VERSION    LABCAL  891208

DISK FILE = 9c88nv.cal

VARIABLE=PRESSURE DOWN

FLUID HEAD OFFSET PRESSURE BIAS =  0.4

12 DATA POINTS          ORDER OF POLY  3

POLY COEF =0.425865E+01 0.998952E-01 0.109585E-07 -.164719E-12

INSTRUMENT      OBSERVED      CALCULATED      DIFFERENCE

1        -30.0000           0.400000       1.261806      -0.861806

2         3403.00         344.919739     344.322235       0.597504

3         6845.00         689.433411     688.501648       0.931763

4         13742.0        1378.442383    1378.659790      -0.217407

5         20627.0        2067.427002    2068.012939      -0.585938

6         27504.0        2756.386963    2756.637451      -0.250488

7         34383.0        3445.322754    3445.213623       0.109131

8         41269.0        4134.234375    4133.918457       0.315918

9         48168.0        4823.121094    4823.025391       0.095703

10         55078.0        5511.983887    5512.005371      -0.021484

11         62004.0        6200.821289    6201.023438      -0.202148

12         65470.0        6545.231445    6545.142578       0.088867

MEAN DEVIATION =  -0.321567E-04

STANDARD DEVIATION =   0.485361E+00

PROGRAM VERSION    LABCAL  891208

DISK FILE = 9c88nv.cal

VARIABLE=PRESSURE UP

12 DATA POINTS          ORDER OF POLY  3

POLY COEF =0.140992E+01 0.997156E-01 0.188603E-07 -.235779E-12

INSTRUMENT      OBSERVED      CALCULATED      DIFFERENCE

1         65470.0        6545.231445    6544.465820       0.765625

2         62006.0        6200.821289    6200.679199       0.142090

3         55094.0        5511.983887    5512.959473      -0.975586

4         48191.0        4823.121094    4824.217285      -1.096191

5         41302.0        4134.234375    4135.424805      -1.190430

6         34422.0        3445.322754    3446.550781      -1.228027

7         27551.0        2756.386963    2758.059570      -1.672607

8         20685.0        2067.427002    2070.010010      -2.583008

9         13816.0        1378.442383    1382.058838      -3.616455

10         6926.00         689.433411     692.866455      -3.433044

11         3460.00         344.919739     346.641907      -1.722168

12        -16.0000           0.400000      -0.185525       0.585525

MEAN DEVIATION =  -0.269592E-04

STANDARD DEVIATION =   0.219664E+01

use order of fit = -1 to obtain DIFFERENCE FILE

Sample of temperature difference file

type 9c88nvt.dat

Temp Lab.      TLAB-Tctd

24.00820      7.2002411E-03

20.01100      7.9994202E-03

15.00620      7.2002411E-03

10.01000      8.0003738E-03

5.025900      8.8996887E-03

3.010100      9.0999603E-03

0.5083000      9.2999935E-03

Sample of PRESSURE difference file

$ type 9c88nvp.dat

Pres. Lab.     PLAB-Pctd

0.4000000       3.400000

344.9197       4.619751

689.4334       4.933411

1378.442       4.242432

2067.427       4.727051

2756.387       5.987061

3445.323       7.022705

4134.234       7.334473

4823.121       6.321289

5511.984       4.184082

6200.821      0.4213867

6545.231      -1.768555

The pressure, temperature, and conductivity difference files can be

plotted by programs such as PLOT5 on the VAX or GRAPHER on the PC. The

drift of a sensor can be visualized, when a sensor variable is plotted

together with previous calibration data.