This file is to document the post post CTD processing of cruise Endeaver 223

by Maggie Cook during the summer,fall 1992.

**********************************************************************

*******                                                        *******

*******                      r/v ENDEAVOR                      *******

*******                       Cruise 223                       *******

*******                                                        *******

**********************************************************************

Chief Scientist:  Bob Pickart

Project Name:  Labrador Sea

Dates (leg1):  22 March 1991 - 06 April 1991

Ports-of-call (leg1):  URI - St. John's

Dates (leg2):  8 April 1991 - 24 April 1991

Ports-of-call (leg2):  St. John's - Iceland

CTD instrument and station numbers:

CTD 10  Stations LEG 1:  1-22

LEG 2: 23-47

CTD 8   Backup instrument not used

-------------------------------------------------------------------------------

Leg 1       CTD Software Tech:   Carol MacMurray, Ruth Curry

Leg 1       CTD Hardware Tech:   Peter Landry

Leg 1             Hydrography:   Rosette samples - Scripps:

nuts               Jim Costello

ox,sa              Dave Muus

Leg 1                            Rosette Freon samples  - Lamont Doherty:

Bill Smethie

Maureen Noonan

Guy Mathieu

Chris Toles

Leg 1                            Rosette Tritium/Helium   - Scot Birdwhistell

Leg 1         URI Marine Tech:   Joe Lewkowicz

Leg 1          Watch Standers:   Amy Ffield, Sandy Fontana

Gwyneth Hufford, Sarah Zimmerman

--------------------------------------------------------------------------------

Leg 2       CTD Software Tech:   Ruth Curry

Leg 2       CTD Hardware Tech:   Peter Landry

Leg 2             Hydrography:   Rosette samples - Scripps:

nuts               Jim Costello

ox,sa              Dave Muus

Leg 2                            Rosette Freon samples  - Lamont Doherty:

Bill Smethie

Maureen Noonan

Guy Mathieu

Chris Toles

Leg 2                            Rosette Tritium/Helium   - Scot Birdwhistell

Leg 2         URI Marine Tech:   Joe Lewkowicz

Leg 2          Watch Standers:   Amy Ffield, Sandy Fontana

Gwyneth Hufford, Sarah Zimmerman

________________________________________________________________________________

Shorebased processor:  Maggie Cook

July - November 1992

MicroVAX Data subdirectory:

BONZO::PRO5:<CTD.EN223D040...>    final data directory

<.CAL>             laboratory calibrations

<.ROS>             files resulting from PC conversion

WRW files to BT2 files

<.HYDRO>           salt and nutrient rosette files

created at sea.  I think T. McKee

has received more updated versions

direct from scripps.

<.HYDROPROGS>      programs used at sea for scaling

rosette data which caused a multitude

of problems.

<.SEA>             files involved in creation of final

SEA files on the PC

<.PCOXY>           input and output files used during

PC computer processing of oxygen

calibration groups

<.MVER>            pressure averaged data files before

any oxygen transformations or editting

--------------------------------------------------------------------------------

************************************************************************

*                                                                      *

*                   CRUISE SUMMARY STATION LISTING                     *

*                                                                      *

************************************************************************

SH CRU STAT DV CST CTD DA MO YR  ST GMT  END GMT  LATITUDE   LONGITUDE   P MAX

-- --- ---- -- --- --- -- -- --  ------  -------  --------   ---------   -----

EN 223   2   E 001 10  28  3 91    317     649    42 23.66   -55  1.28   3799.

EN 223   3   E 001 10  28  3 91   1248    1438    42 47.71   -55  0.40   4477.

EN 223   4   E 001 10  28  3 91   2035    2226    43 10.12   -55  0.05   4305.

EN 223   5   E 001 10  29  3 91    242     407    43 27.10   -54 59.90   4051.

EN 223   6   E 001 10  29  3 91    918    1103    43 39.80   -54 47.82   3765.

EN 223   7   E 001 10  29  3 91   1533    1650    43 54.14   -54 38.00   3415.

EN 223   8   E 001 10  29  3 91   2130    2238    44  7.83   -54 28.15   2689.

EN 223   9   E 001 10  30  3 91    303     404    44 20.65   -54  8.00   2305.

EN 223  10   E 001 10  30  3 91    854    1013    44 34.00   -54  3.01   1311.

EN 223  11   E 001 10   1  4 91   1355    1542    43 53.50   -42  9.33   4853.

EN 223  12   E 001 10   1  4 91   2035    2221    43 58.00   -42 28.10   4875.

EN 223  13   E 001 10   2  4 91    233     414    44  2.40   -42 50.00   4877.

EN 223  14   E 001 10   2  4 91   1008    1148    44 28.02   -43 28.02   4873.

EN 223  15   E 001 10   2  4 91   1854    2038    44 20.17   -44-21.05   4803.

EN 223  16   E 002 10   3  4 91    835    1015    44 30.15   -45-20.90   4235.

EN 223  17   E 001 10   4  4 91   1742    1926    44 41.92   -46 -4.94   3747.

EN 223  18   E 001 10   4  4 91    120     232    44 54.96   -46-36.82   3609.

EN 223  19   E 001 10   4  4 91    648     844    45  8.94   -46-56.13   3429.

EN 223  20   E 001 10   4  4 91   1325    1439    45 20.28   -47-14.60   2837.

EN 223  21   E 001 10   4  4 91   2137    2240    45 29.93   -47-38.27   2003.

EN 223  22   E 001 10   5  4 91    342     504    45 40.60   -47-51.65   1255.

EN 223  23   E 001 10  10  4 91    133     215    47  4.96   -47 -2.93   1103.

EN 223  24   E 001 10  10  4 91    518     605    47  5.20   -46-40.50   1069.

EN 223  25   E 001 10  10  4 91   1130    1143    47 10.16   -45-44.05    265.

EN 223  26   E 001 10  10  4 91   1518    1533    47 29.06   -45 -9.04    219.

EN 223  27   E 001 10  10  4 91   1833    1849    47 52.04   -44 47.91    265.

EN 223  28   E 001 10  10  4 91   2218    2251    48 22.05   -44 32.17    899.

EN 223  29   E 001 10  11  4 91    218     312    48 38.20   -44 23.27   1619.

EN 223  30   E 001 10  11  4 91    715     816    48 57.03   -44 10.96   2401.

EN 223  31   E 001 10  11  4 91   1134    1255    49  8.06   -43 59.98   3249.

EN 223  32   E 001 10  11  4 91   1635    1810    49 19.96   -43 49.17   3951.

EN 223  33   E 001 10  12  4 91     23     214    49 30.45   -43 37.65   4035.

EN 223  34   E 001 10  12  4 91   1157    1402    49 44.95   -43 10.25   4263.

EN 223  35   E 001 10  13  4 91   1413    1643    50  4.08   -42 58.44   4337.

EN 223  36   Y 001 10  13  4 91   2313      43    50 22.89   -42 48.09   4323.

EN 223  37   Y 001 10  17  4 91    448     527    48 42.51   -44 13.02   1001.

EN 223  38   Y 001 10  17  4 91    800     842    48 47.21   -44  7.45   1007.

EN 223  39   Y 001 10  17  4 91   1059    1133    48 46.41   -44 17.83    999.

EN 223  40   Y 001 10  17  4 91   1330    1409    48 38.19   -44 16.55    999.

EN 223  41   Y 001 10  17  4 91   1625    1703    48 40.00   -44  7.14    999.

EN 223  42   Y 001 10  19  4 91    330     448    53 57.01   -46 25.79   3653.

EN 223  43   Y 001 10  19  4 91   1344    1512    55 12.02   -47  8.06   3623.

EN 223  44   Y 001 10  20  4 91   1419    1522    58 18.57   -42 38.35   2951.

EN 223  45   Y 001 10  20  4 91   1825    1925    58 30.06   -42 41.09   2647.

EN 223  46   Y 001 10  20  4 91   2218    2309    58 39.82   -42 48.03   2361.

EN 223  47   Y 001 10  21  4 91    433     552    58  4.94   -42 29.87   3161.

CRUISE     MIN/MAX DATE    MIN LAT/LON MX LAT/LON   INST   NV MIN MAX TOT P WH

--------- ----------------- ----------- ---------- -------- -- --- --- --- - --

EN223D040 91- 3-28 91- 4-21   42.  -55.  59.  -42. 10 0 0 0  5   2  47  46 Y

************************************************************************

*                                                                      *

*                     A C Q U I S I T I O N                            *

*                                                                      *

************************************************************************

Data was acquired using the MicroVAX CTD acquisition program developed

by Julie Allen.  Original at sea processing was done on the

MicroVAX using the MLTRG to PRSORT sequence for processing salts and

the OXFITED to OXCALCL sequence for processing oxygens.  A new suite of

R. Millard CTD processing programs had been developed prior to this

cruise for use on both the PC and the VAX.  However, only a few

test stations were processed this way during the cruise.  The

MLTRG and PRSORT programs used were the recently alterred 1991 versions

(MLTRG91, PRSORT91) which incorporated a pressure temperature variable

into the scaling algorithm to accomodate the new titanium pressure

transducer.  It also encorporated the SAL78 subroutine which was corrected

by replacing C1535 with 42.914  (formerly 42.909).

The following version of the MicroVAX acquisition was used to collect

the data onto a VAXSTATION 3600 (URI Shipboard workstation):

WHOI AQUI89 acquistion package (Version 1.0)

NOTE:

There was a test station #999 which showed multiple integer*2 errors during

the cast.  It appears that integer*2 in the AQUI89 program was changed to

integer*4 in an attempt to correct the problem before station #1 (This is

a guess).  Station #1 was collected with a clean trace but the resulting

file was in a format unreadable by any post processing programs.  I believe

that integer*4 was set back to integer*2 before collecting the remainder of

the stations (This is also a guess).  There is nothing to suggest that any

attempt to salvage the unreadable station #1 data was made (I would have

thought that an attempt to reread the audio data might have been done at

sea if station # 1 was important enough to warrant the effort).  There is

rosette sample data for station # 1.

Data was logged to    a.  Audio tape

b.  9T tape

c.  VaxStation 3600 file (###A###.RAW) in CTD78 format

(*.WRW,*.WSC,*.HED,*.ERR) in ASCII format

*.RAW files were backed up on TK50 tape:

See 'CTD process checklist' notebook for specifics:

AQRAW1 :   stations test 999, 1 (unreadable)

stations 02 - 10

AQRAW2 :   stations 11 - 22

AQRAW3 :   stations 23 - 36

AQRAW4 :   stations 35 - 47

The at sea COMMON used for data acquisition for CTD #10 follows:

ID   attribute 1     attribute 2     slope        bias           sens lag

PR   0.244853E-08  -0.518678E-13   0.100212E+00  -0.134196E+00   0.000000E+00

TE   0.184585E-11   0.000000E+00   0.499957E-03  -0.402169E-02   0.250000E+00

CO  -0.650000E-05   0.150000E-07   0.100456E-02  -0.143578E-01   0.000000E+00

SW   0.280000E+01   0.300000E+04   0.100000E+01   0.000000E+00   0.000000E+00

OC  -0.360000E-01   0.115000E-03   0.148000E-02   0.000000E+00   0.000000E+00

OT   0.750000E+00   0.000000E+00   0.128000E+00   0.000000E+00   0.000000E+00

TP   0.000000E+00   0.000000E+00   0.100000E+01   0.000000E+00   0.000000E+00

RT   0.000000E+00   0.000000E+00   0.000500E+00  -2.000000E+00   0.000000E+00

QU   0.000000E+00   0.000000E+00   0.100000E+01   0.000000E+00   0.000000E+00

This common uses precruise calibrations for pressure, temperature and

conductivity.  The calibration information for these are in the following files:

LABCAL,RUSKCAL (PC based programs)

CTD 10    [CTD.EN223D040.CALS]10MR91TC.CAL   main instrument  TE,CO

[CTD.EN223D040.CALS]10MR91PR.CAL                    PR

CTD  8    [CTD.EN223D040.CALS]8MR91TC.CAL   backup instrument TE,CO

[CTD.EN223D040.CALS]8MR91PR.CAL                     PR

note:     CTD #8 was not used on this cruise.

The computer used for shipboard processing of the CTD data was a VAXstation 3600

(DRAGON) belonging to URI.  It had 16 mb of memory and 1 gigabyte of disk space.

Downtrace CTD data was editted using the MicroVAX version of the HP

CTDED program converted for VAX use by Mary Hunt.

Input:      MicroVAX Acquisition CTD78 format raw data 9T tapes.

Output:     MicroVAX disk files [CTD.EN223D040.CTDED78]####D###.EDT

Program CTLOG was used for copying the above editted CTD78 format VAX

data files  to 9T tape: (See 'CTD process checklist' notebook for specifics)

LS01   stations    2 -  9

LS02   stations   10 - 22

LS03   stations   23 - 36

LS04   stations   37 - 47

TK50 VMS backup tapes were also made of these *.EDT files:

See 'CTD process checklist' notebook for specifics:

CTDED1  :  stations 02 - 10

CTDED2  :  stations 11 - 28

CTDED3  :  stations 29 - 36

The following program for initial creation of VAX format water sample files was

used:

Program:  WS_EN223

THIS IS THE PROGRAM WHICH SCALED PR AND TE IN THE "RAW" DATA FILE!!

WATER SAMPLE FILES OUTPUT BY THIS PROGRAM WERE USED AT SEA AND INITIALLY POST

CRUISE TO CALIBRATE THE CTD DATA ... THEY INTRODUCED A .002 ERROR IN THE

DEEP WATER SALTS.

Water sample filenames:   Salinity and Oxygen Rosette Samples

En223c10s.mlt - all stations

En223c10s.DYN

THESE FILES HAVE DOUBLY SCALED PR AND TE!  THEY SHOULD NOT BE USED!  THE *.MLT

FILE WAS, HOWEVER, USED TO FOR GRABBING THE ROSETTE HYDRO DATA AND MERGING IT

INTO A FILE WITH RAW PRESSURE AND TEMPERATURE DATA FOR CALIBRATION WORK ON THE

PC.

NUTRIENT WATER SAMPLE FILENAMES:

PRO5:[CTD.EN223D040.HYDRO] Sta#.HYD

Terry McKee has merged the nutrient data into the SEA file.  She used a more

current version of the data set from SCRIPPS rather than the *.HYD files

contained in a subdirectory here.

There is a listing of changes made to the original rosette water samples in

terms of misfires, etc.  The file name is HYDROEDIT.DOC.  This may be

irrelevent now that T. McKee has merged nutrient data direct from SCRIPPS.

*******************************************************************************

CTD data were pressure averaged at sea using the VAX PRSORT91 program

which encorporates a pressure temperature variable into the scaling of

the new titanium pressure transducer used on this cruise.  However, there

was no calibration done on these instruments for pressure temperature,

so pressure temperature adjustments were not applied.  This is not

expected to be a problem since the data was collected in the North

Atlantic where temperature variations surface to depth are minimal.

Data AT SEA were pressure averaged with PRE-CRUISE CUBIC pressure

calibrations and PRE-CRUISE QUADRATIC temperature calibrations.

The FINAL POST-CRUISE version of the data were pressure averaged with

PRE-CRUISE QUADRATIC pressure calibrations and POST-CRUISE QUADRATIC

temperature calibrations.

The conductivity calibrations used for pressure averaging at sea were

derived by doing best fits to groups of stations and their respective

rosette data using program MLTRG91.  This was believed to be a very

clean, stable and well calibrated data set at the time.  However, at

the time of post cruise processing it would be discovered that there

was a distinct problem (introduced during the at sea processing of the

rosette water sample data) with the method of deriving calibrations

for the conductivity sensor for this cruise.  Upon close observation,

it was noted that there was a distinct and consistent .002 discrepancy

between deep water CTD and rosette water sample data.  This difference

appeared consistently even though program MLTRG91 showed uniform and

acceptable histograms of fits of the CTD and rosette water sample data.

After much scrutiny of program MLTRG91 (since it had been recently

modified), as well as questions regarding the CTD and/or the rosette

data, it was finally discovered that the (presumeably) RAW rosette data

file being used to fit the data for conductivity calibrations had in

fact been  scaled already with precruise pressure and temperature

laboratory calibrations.  Thus, through program MLTRG91 the scaled

pressure and temperature data were being doubly scaled, and those doubly

scaled values were being used to invert rosette salinities to obtain

rosette conductivities for comparison to CTD conductivities.  To

backtrack through the loop to retrieve the original unscaled data would

be a cumbersome, confusing prospect.

To make a long story short, Bob Millard convinced us that the  best

method for overcoming all this was to step back and revert to

processing of the water sample data on the new "streamlined" PC CTD

processing system.  Input would be the unadulterated raw rosette water

sample files created by the CTD acquisition program.  This, as opposed

to using a water sample file which was slotted to be "raw" but in fact

had been manipulated in terms of calibrations.  (The logic here

would circumvent the fear that perhaps  other manipulations had been

made to that presumed 'raw' data file that we did not know about.)

Of course, the sequence of processing events diverts here from "normal"

since this cruise was not slated to be processed on the PC, and so many

alternatives to the path for PC processing needed to be adapted for this

particular cruise.

****************************************************************************

POST CRUISE PC PROCESSING OF ENDEAVOR CRUISE 223 CTD DATA

****************************************************************************

PC data processing went as follows (with many hiccups along the way!)

1.  Program BTLFMTVX was run on the PC

a.  Input:   0###A00#.WRW   rosette file created by VAX acquisition

program AQUI89

0###A00#.C00   calibration file for that station

(this created using edittor to represent

file which would have been created by

PC acquisition program).  It contains

pre-cruise calibrations for PR,TE,CO and

nominal cals for OC,OT.  Some of the comment

lines are inconsistent with the truth, but

these are consistent with the format of all

*.C00 files used on cruises thus far.  I

decided that minimal changes would introduce

minimal problems?

0###A00#.HED   header file for that station

(This also created using edittor to

represent file which would have been

created by PC acquisition).

b.  Output:  0###A00#.BT2   Rosette file for each station which

contains columns with -99.0000 for input of

rosette salt and oxygen data.

0###A00#.C00

2.  At this point all stations of the cruise were merged together to form one

file, ALLSTA.FNL .

Then, file C10S.MLT was brought over to the PC from the VAX.  This file

was created on the VAX using MLTRG91.  It has rosette salts and oxygens

in it as well as CTD pressure and temperature.  However, as noted earlier,

the pressure and temperature had been scaled incorrectly at sea.  This MLT

file is used here only to extract the rosette salts and oxygens for all

stations and input them into the ALLSTA.FNL file.  Utility PW on the

PC was used to merge rosette salts and oxygen values into the ALLSTA.FNL

file (column by column).

For use later on, the ALLSTA.FNL file (which now has all stations merged

together, includes rosette salts and oxygens, and is the same format as a

0###A00#.BT2 file), is divided back up station by station into 0###S00#.BT2

files.

2a. Alternatively, on a cruise acquired and worked up on the PC alone,

a template of each station from step one in the above process would

be given to each hydrographic group to merge in their rosette results.

For example the WHOI hydro group would run their program HYDMERGE to input

rosette salts and oxygens and pass back a *.SOX file to the CTD data

processor.  Other groups would create *.NUT and *.CFC files.  Then program

SEAMERGE would be run by the CTD data processore which would merge all the

rosette data groups (SOX,NUT,CFC,...) into one file: *.SEA.

3a. So now we have a ALLSTA.FNL file which is the same format as a *.BT2 file

but has all stations merged together and includes rosette salts and oxygens

instead of just -99.0000 default values.  The norton edittor was then used

to add a header to this file (the format of which conductivity fitting

programs CONVERT and CONCALD expect for file input).  The file is now

called EN223BTL.LST .

4a. File EN223BTL.LST is copied to four files names on the PC, each of which

will act as input files to program CONVERT.  There will be four station

groups for fitting salts.  These four groups were previously determined

during multiple passes of running MLTRG91 before discovering that there

was an inherent problem with the raw data being input to that program.

(The four groups are pointed out on the delta conductivity plot created

with grapher after doing a CONVERT and CONCALD fit to the entire water

sample file, EN223BTL.LST .)

EN223BTL.LST was copied to EN223A.LST, *B.LST, *C.LST, *D.LST.  Each of

these A,B,C,D files were editted to include only the stations that would

be used to fit for conductivity for that group.

EN223A.LST     stations 02-09 (except 8)

EN223B.LST     stations 11-20

EN223C.LST     stations 31-36 (except 34)

EN223D.LST     stations 43-47

Then, each of these above files were copied to files of the same name, but

with *AA.LST,*BB.LST,*CC.LST,*DD.LST, for more refined processing of the

conductivity fitting procedure as is described below.  This naming of one

file to multiple names is due to the fact that these programs on the PC

expect all input and output file names to be identical. In order to

not overwrite necessary initial processing procedures and output results

with the more refined ones, then we must have these multiple file names.

Also, it was necessary to create *.C00 files to correspond to each of the

above file names.  These *.C00 files are of the EG&G format and contain

the pre-cruise quadratic pressure and post cruise quadratic temperature

cals that will be used in the final data load.

So...

5.  Run PC program CONVERT (C:\CTD92\CAL\CONVERT\CONVERT)

a.  Input:    EN223A.LST  (B,C,D,AA,BB,CC,DD also)

EN223A.C00  (B,C,D,AA,BB,CC,DD also)

EN223A.PBI  (B,C,D,AA,BB,CC,DD also)

this file contains the on deck pre-cast

pressure bias reading for each station.

b.  Output:   EN223A.WSD  (B,C,D,AA,BB,CC,DD also)

This program CONVERT is formatting the rosette file for input to program

CONCALD which does the actual fitting of the CTD data to the rosette salt

data.

6.   Run PC program CONCALD (C:\CTD92\CAL\CONCALD\CONCALD)

1.  At this stage, program CONCALD was used to fit for conductivity bias

and slope over the entire water column (0-6500 db)

a.   Input:   EN223A.WSD  (B,C,D also)

b.   Output:  EN223A.HIS  (B,C,D also)  histogram of residuals to fit

EN223A.REJ  (B,C,D also)  rejected observations

EN223A.RES  (B,C,D also)  residuals to fit

(for plotting with grapher!)

2.  Next, program CONCALD was used to fit for slope alone in the deep

water (2500-6500 db) having removed the conductivity bias term

(determined above) from the data.

a.  Input:    EN223AA.WSD (BB,CC,DD also)

b.  Output:   EN223AA.HIS (BB,CC,DD also)

EN223AA.REJ (BB,CC,DD also)

EN223AA.RES (BB,CC,DD also)

RESULTS OF STEP 6 CONDUCTIVITY FITS:  (***** columns used for final load)

A,AA  applied to stations 02-10

B,BB  applied to stations 11-22

C,CC  applied to stations 23-40

D,DD  applied to stations 41-47

Step 1: CO Bias     Step 1: CO Slope        Step 2: Refined CO Slope

***************     ----------------        ************************

EN223A     .41809092E-2         .10034343E-2          EN223AA  .10034298E-2

EN223B     .66004314E-2         .10034158E-2          EN223BB  .10034147E-2

EN223C     .21887135E-1         .10028947E-2          EN223CC  .10028896E-2

EN223D     .51382207E-3         .10033635E-2          EN223DD  .10033625E-2

7.  Program PRSORT91 was used on the VAX to pressure average the data using

Pre cruise Pressure calibrations, Post cruise temperature calibrations,

and the above conductivity calibrations as well as nominal oxygen

calibrations which would be important to note later in the processing of

the oxygen data.  (Note: that this PRSORT91 step had been done numerous

times initially while trying to decipher the problem with doing processing

on the VAX instead of the PC).

Pre Cruise Down/Up Quadratic Pressure Calibrations were used to PRSORT the

final data using -(pre-lowering deck unit reading) for the PR bias term.

Note that this is a combination of the down and the up trace used for the

fit.  The new titanium pressure transducer presumeably has none of the

pressure hystoresis characteristics seen in earlier transducers.

BIAS        SLOPE         ATTR 1

PRE-CRUISE  CTD10  -.477829E0   .100306E+0   -.219531E-8

Post cruise calibrations were used for final data load.

This is based on the fact there there was a 4 millidegree shift in the

temperature sensor calibration pre cruise to post cruise.  The start of

this cruise was extremely rough, and the CTD slammed into the side of the

ship's hull several times.  We are speculating that the large temperature

calibration shift occurred during these first two casts.

BIAS         SLOPE         ATTR 1       TE time Lag

POST-CRUISE CTD10  -.614282E-2   .499557E-3    .785609E-11     0.25 seconds

7a. The scale factor records used to pressure average the 4 conductivity station

groups listed above are as follows.  The pressure bias term changes station

to station depending on the on deck pre-cast-lowering deck unit reading:

Stations 02-10:

ID    ATTR. 1     SLOPE      ENG. BIAS   T.CONST.     ATTR. 2

PR-0.219531E-08 0.100306E+00 0.110000E+01 0.000000E+00 0.000000E+00

TE 0.785609E-11 0.499557E-03-0.614282E-02 0.250000E+00 0.000000E+00

CO-0.650000E-05 0.100343E-02 0.418091E-02 0.000000E+00 0.150000E-07

SW 0.280000E+01 0.100000E+01 0.000000E+00 0.000000E+00 0.300000E+04

OC-0.360000E-01 0.148000E-02 0.000000E+00 0.000000E+00 0.115000E-03

OT 0.500000E+00 0.128000E+00 0.000000E+00 0.000000E+00 0.000000E+00

TP 0.000000E+00 0.100000E+01 0.000000E+00 0.000000E+00 0.000000E+00

RT 0.000000E+00 0.500000E-03-0.200000E+01 0.000000E+00 0.000000E+00

QU 0.000000E+00 0.100000E+01 0.000000E+00 0.000000E+00 0.000000E+00

Stations 11-22:

ID    ATTR. 1     SLOPE      ENG. BIAS   T.CONST.     ATTR. 2

PR-0.219531E-08 0.100306E+00-0.200000E+00 0.000000E+00 0.000000E+00

TE 0.785609E-11 0.499557E-03-0.614282E-02 0.250000E+00 0.000000E+00

CO-0.650000E-05 0.100341E-02 0.660043E-02 0.000000E+00 0.150000E-07

SW 0.280000E+01 0.100000E+01 0.000000E+00 0.000000E+00 0.300000E+04

OC-0.360000E-01 0.148000E-02 0.000000E+00 0.000000E+00 0.115000E-03

OT 0.500000E+00 0.128000E+00 0.000000E+00 0.000000E+00 0.000000E+00

TP 0.000000E+00 0.100000E+01 0.000000E+00 0.000000E+00 0.000000E+00

RT 0.000000E+00 0.500000E-03-0.200000E+01 0.000000E+00 0.000000E+00

QU 0.000000E+00 0.100000E+01 0.000000E+00 0.000000E+00 0.000000E+00

Stations 23-40:

ID    ATTR. 1     SLOPE      ENG. BIAS   T.CONST.     ATTR. 2

PR-0.219531E-08 0.100306E+00 0.600000E+00 0.000000E+00 0.000000E+00

TE 0.785609E-11 0.499557E-03-0.614282E-02 0.250000E+00 0.000000E+00

CO-0.650000E-05 0.100289E-02 0.218871E-01 0.000000E+00 0.150000E-07

SW 0.280000E+01 0.100000E+01 0.000000E+00 0.000000E+00 0.300000E+04

OC-0.360000E-01 0.148000E-02 0.000000E+00 0.000000E+00 0.115000E-03

OT 0.500000E+00 0.128000E+00 0.000000E+00 0.000000E+00 0.000000E+00

TP 0.000000E+00 0.100000E+01 0.000000E+00 0.000000E+00 0.000000E+00

RT 0.000000E+00 0.500000E-03-0.200000E+01 0.000000E+00 0.000000E+00

QU 0.000000E+00 0.100000E+01 0.000000E+00 0.000000E+00 0.000000E+00

Stations 41-47:

ID    ATTR. 1     SLOPE      ENG. BIAS   T.CONST.     ATTR. 2

PR-0.219531E-08 0.100306E+00-0.500000E+00 0.000000E+00 0.000000E+00

TE 0.785609E-11 0.499557E-03-0.614282E-02 0.250000E+00 0.000000E+00

CO-0.650000E-05 0.100336E-02 0.513822E-03 0.000000E+00 0.150000E-07

SW 0.280000E+01 0.100000E+01 0.000000E+00 0.000000E+00 0.300000E+04

OC-0.360000E-01 0.148000E-02 0.000000E+00 0.000000E+00 0.115000E-03

OT 0.500000E+00 0.128000E+00 0.000000E+00 0.000000E+00 0.000000E+00

TP 0.000000E+00 0.100000E+01 0.000000E+00 0.000000E+00 0.000000E+00

RT 0.000000E+00 0.500000E-03-0.200000E+01 0.000000E+00 0.000000E+00

QU 0.000000E+00 0.100000E+01 0.000000E+00 0.000000E+00 0.000000E+00

8.  Next it was necessary to take the rosette file we had been working with on

the PC and convert it to a format formidable to doing overplots of Theta/S

on the vax  (i.e.  Make a *.DYN file).

a.  PC program CONVERT must be rerun inputting the aggregate of all stations

in file EN223BTL.LST.  Associated *.C00 (containing the best pressure

and temperature calibrations) and the  *.PBI file (containing pre-cast

on deck deck unit PR readings) must be available for input.  Output is

EN223BTL.WSD .

b.  File EN223BTL.WSD is transferred to the VAX.

c.  Program CTDA:[CTDEV.WS.HYDOUTV]HYDOUTV is run.

1.  Input:   EN223BTL.WSD  (had to be in [CTD.EN223D040] directory)

Output:  EN223BTL.DYN  (hydro file formatted for CTD80D plotting on

the VAX.  Can be listed with program

LISTGRID)

9.  Also, it was necessary to take the rosette file we had been working with on

the PC and convert it to a WOCE format hydro water sample file for use by

Thierry McKee.  Of course, this would need to be revised later on when

oxygen calibrations had been finalized to update the CTD oxygen data in the

file.

a.  Take the *S*.BT2 files.  (These are basically the ALLSTA.FNL file

divided up station by station.  They are

also equivalent to the *A*.BT2 files which

have real rosette data merged into the -99

columns for salts and oxygens)

b.  Create an EG&G format *.C00 file for EACH STATION containing the

best calibrations available. (At this point we have PR,TE,CO).

c.  Copy a TEMPLATE.HDR file into your directory and make sure it includes

information pertinent to this cruise.  (i.e. EN223 cruise dates)

d.  Run PC program WOCTMPV1 on each station.  This will rescale the rosette

file with the updated calibrations and output a *.CTP file from the

input *S*.BT2 file.

OXYGEN FITS:

1.  The DYN file created above was run through VAX program

CTDA:[CTDEV.CAL.OXEXTVX]OXEXTVX1

to create EN223BTL.OXY.  This file would be used for fitting oxygens on the

PC the same way that MLTOXDWN output files would be used for fitting oxygens

on the VAX.

2.  The EN223BTL.OXY file was brought over to the PC for fitting oxygens.

Program OXFITMR was run on the entire cruise worth of data inputting this

file and outputting histograms and delta O2 files for plotting.  It must

be noted that OXFITMR assumes an OC slope of .0014, so the OC slope of

.00148 used during pressure averaging needed to be specified whenever

running the program in order to obtain valid results.

Run Program OXFITMR

input:   *.OXY   containing only stations included in fit

output:  *.DAT   for plotting delta's etc.

*.HIS   histogram of fitting results

*.REJ   scans rejected from the fit

FOR020.DAT   rename this to *.CLC

The *.CLC file is used on the VAX with program OXCALCL

for transforming the oxygens.

3.  Many station groups for fitting were chosen based on plots of the above

delta O2's.  Program OXFITMR was run on many sets of various station

groupings (inputting editted versions of the *.OXY file containing only

the stations of each group).  To make another long story short, it was

discovered (after many checks of the data and of the functionality of

the fitting programs) that the oxygen sensor used through station 35

was not performing properly.  In fact, it appears that the calibrations

required to describe sensor characteristics varied even within a data

trace.  After many, many tests it was decided to zero out all oxygen data

for stations 2-35 and depend entirely on the rosette oxygen data for those

stations.  The remainder of the cruise had well behaved oxygen traces

which were collected with a different oxygen sensor.

4.  The following calibration results, although fits to small station groups,

still failed to describe the oxygen sensor properly.  These stations,

station 2-35, had all CTD oxygens ultimately set to zero.

-0.037  0.1058E+01  0.1146E-03 -0.3710E-01  0.5053E+00  0.4312E+01/  sta 03-05

0.126  0.8828E+00  0.8167E-04 -0.3803E-01  0.9352E+00  0.8000E+01/  sta 06-07

0.019  0.9481E+00  0.1233E-03 -0.3122E-01  0.5200E+00  0.3286E+01/  sta 08-09

-0.072  0.1003E+01  0.1287E-03 -0.3343E-01  0.4438E+00  0.8000E+01/  sta 11-12

-0.017  0.9821E+00  0.1118E-03 -0.3575E-01  0.5381E+00  0.8000E+01/sta 13,15,17

-0.019  0.9857E+00  0.1201E-03 -0.3675E-01  0.7870E+00  0.8000E+01/  sta 14,16

0.076  0.8940E+00  0.9919E-04 -0.2859E-01  0.3489E+00  0.8000E+01/  sta 19-21

0.011  0.9094E+00  0.8467E-04 -0.2844E-01  0.4548E+00  0.4729E+01/  sta 23-27

0.028  0.9215E+00  0.1079E-03 -0.3088E-01  0.6578E+00  0.8000E+01/  sta 28-29

0.102  0.8195E+00  0.6773E-04 -0.5819E-02  0.6500E+00  0.8000E+01/  sta 30-32

-0.091  0.9358E+00  0.1576E-03 -0.3324E-02  0.6500E+00  0.8000E+01/  sta 33-34

-0.014  0.9468E+00  0.8820E-04 -0.1622E-01  0.6500E+00  0.8000E+01/  sta 35 only

The new oxygen sensor proved to be stable and calibrations derived using

program OXFITMR allowed CTD traces to fit rosette water sample data very

well in  the Theta profile.

0.034  0.8367E+00  0.1383E-03 -0.3612E-01  0.7263E+00  0.5283E+01/  sta 36-38

0.060  0.9550E+00  0.1566E-03 -0.6018E-01  0.4200E+00  0.4251E+01/  sta 39-41

0.009  0.9344E+00  0.1446E-03 -0.2990E-01  0.4237E+00  0.4000E+01/  sta 42-47

FINAL H2O ROSETTE FILE:

1.  Program OXUPDATE must be run to put downtrace oxygens in the final

form of the rosette file.

Run Program OXUPDATE

input:  ___S___.BT2

___S___.C00

EN223BTL.OXY

output:  ___S___.BT3

2.  Program WOCTEMPV1 must be rerun to do the final scaling of salts and

oxygens to the rosette file.  You must have *.C00 files with correct

conductivity and oxygen scaling factors for every station.

input:  ___S___.BT3

output: ___S___.CTP

3.  Merge all the final *S*.CTP files together to create the final *.SEA file.

**************************************************

*********                                *********

*********        QUALITY CONTROL         *********

*********                                *********

**************************************************

<CTD.CTDATA.EXE>CTDED.EXE  (27-oct-86)           EN223CTDED.COM

To correct erroneous data scans (i.e. interpolate over pressure

bounds to get rid of salinity and oxygen spikes in the data)

Outputs 'E' version.

Runs CTDED to correct .CTD data files.

Branch 19 to interpolate, VAR 1=temp, 2=salinity, 3=oxygen.

Branch 20 to set data value of first scan for OX to -9 (since

Branch 25 doesn't seem to deal with 1st data scan properly).

Branch 25 to set entire trace worth of oxygens to -9 for first 35 stations.

Branch 23 to average noisy salt data station 34

Branch 22 to apply bias to entire profile salt data station 34

Correcting data spikes on several stations!

Zeroing out oxygens on stations 2-35.  It was determined during

post cruise processing that the oxygen sensor was out to lunch until

the sensor was replaced after station 35.  It appears that the calibrations

for the sensor were variable even within a cast until station 35!

STATION       VARIABLE       ACTION        P1, P2      COMMENTS

------------     --------    -----------     ---------    --------------------

2             SA           19/         3625,3633    INTERPOLATE

18             SA           19/         2107,2111    INTERPOLATE

22             SA           19/          903, 907    INTERPOLATE

34             SA           23/          1,6000      AVERAGE OVER 10 DB

34             SA           22/          1,6000      APPLY +.0055 BIAS

CTD OXYGEN DATA WAS ALL REPLACED WITH -9.000 FOR STATIONS 2-35 DUE TO THE

UNRULY BEHAVIOR OF THE OXYGEN CELL.

_______________________________________________________________________________

<CTD.CTDATA.EXE>CRSUMARY1.EXE   (12-jan-83)

To create SUBINDEX.CTD file of selected versions of CTD data files

FINAL ARCHIVED VERSIONS OF SUBINDEX.CTD

DATES   DECEMBER 1992        SUBINDEX.CTD   final YE version data files

...............................................................................

<CTD.CTDATA.EXE>INDEXED.EXE   (21-aug-84)

To correct erroneous header information in SUBINDEX.CTD files

and transfer that corrected info to the station data files.

branch 3:  auto-check mode to scan for errors, run interactively

branch 1:  to correct header info in SUBINDEX.CTD file

branch 2:  transfer header info

from SUBINDEX.CTD file to data files

Header information corrections:

Command file:    EN223INDEXED.COM

Stations 2-14 had their header information alterred to correct for longitude

having been originally input as east (+) instead of west (-) .

Station 24 had an erronious end position which was corrected here.

**************************************************

*********                                *********

*********           B A C K U P          *********

*********                                *********

**************************************************

DATE         TAPE LABEL         WHERE STORED      SAVE_SET NAME(S)

------       ------------       --------------     ----------------

07-DEC-1992       EN223 (9T)       CLARK 349        1. EN223D040.BCK

PRO5:[CTD.EN223D040...]

07-DEC-1992       EN223 (TK50)     CLARK 349        1. EN223D040.BCK

PRO5:[CTD.EN223D040...]

07-DEC-1992       EN223  (9T)      CLARK ANNEX      1. EN223D040.BCK

PRO5:[CTD.EN223D040...]

07-DEC-1992       EN223  (9T)      BOB PICKART      1. EN223D040.BCK

PRO5:[CTD.EN223D040...]

**************************************************

*********                                *********

*********             N O D C            *********

*********                                *********

**************************************************

NODCEXCH tapes have not yet been made of the final EN223 data set.

<CTD.CTDATA.EXE>NODCEXCH.EXE  (29-may-85)

WHERE       QUALITY CHECK

DATE         ALL STAS?    DENSITY      STORED      G HEIMERDINGER

----         ---------    -------      ------      --------------

December 7,1992    Y          1600        Clark 349

Clark Annex

Bob Pickard