**********************************************************************

*******                                                        *******

*******                      r/v ENDEAVOR                      *******

*******                       Cruise 223                       *******

*******                                                        *******

**********************************************************************

Chief Scientist:  Bob Pickart

Project Name:  Labrador Sea

Dates (leg1):  22 March 1991 - 06 April 1991

Ports-of-call (leg1):  URI - St. John's

Dates (leg2):  8 April 1991 - 24 April 1991

Ports-of-call (leg2):  St. John's - Iceland

CTD instrument and station numbers:

CTD 10  Stations LEG 1:  1-22

LEG 2: 23-47

CTD 8   Backup instrument not used

Aborted stations:

-------------------------------------------------------------------------------

Leg 1       CTD Software Tech:   Carol MacMurray, Ruth Curry

Leg 1       CTD Hardware Tech:   Peter Landry

Leg 1             Hydrography:   Rosette samples - Scripps:

nuts               Jim Costello

ox,sa              Dave Muus

Leg 1                            Rosette Freon samples  - Lamont Doherty:

Bill Smethie

Maureen Noonan

Guy Mathieu

Chris Toles

Leg 1                            Rosette Tritium/Helium   - Scot Birdwhistell

Leg 1         URI Marine Tech:   Joe Lewkowicz

Leg 1          Watch Standers:   Amy Ffield, Sandy Fontana

Gwyneth Hufford, Sarah Zimmerman

--------------------------------------------------------------------------------

Leg 2       CTD Software Tech:   Ruth Curry

Leg 2       CTD Hardware Tech:   Peter Landry

Leg 2             Hydrography:   Rosette samples - Scripps:

nuts               Jim Costello

ox,sa              Dave Muus

Leg 2                            Rosette Freon samples  - Lamont Doherty:

Bill Smethie

Maureen Noonan

Guy Mathieu

Chris Toles

Leg 2                            Rosette Tritium/Helium   - Scot Birdwhistell

Leg 2         URI Marine Tech:   Joe Lewkowicz

Leg 2          Watch Standers:   Amy Ffield, Sandy Fontana

Gwyneth Hufford, Sarah Zimmerman

________________________________________________________________________________

Shorebased processor:

MicroVAX Data subdirectory: DUA0:<CTD.EN223d040>

--------------------------------------------------------------------------------

************************************************************************

*                                                                      *

*                   CRUISE SUMMARY STATION LISTING                     *

*                                                                      *

************************************************************************

************************************************************************

*                                                                      *

*                     A C Q U I S I T I O N                            *

*                                                                      *

************************************************************************

-----------------------

PRE-CRUISE CALIBRATIONS

-----------------------

LABCAL,RUSKCAL (PC based programs)

CTD 10    10MR91TC.CAL   main instrument  TE,CO

10MR91PR.CAL                    PR

CTD  8    8MR91TC.CAL   backup instrument TE,CO

8MR91PR.CAL                     PR

-----------------------

AT SEA DATA ACQUISITION

-----------------------

VAXSTATION 3600 (URI Shipboard workstations)

with WHOI AQUI89 acquistion package (Version 1.0)

Logging data to:   Audio

9T

VaxStation 3600 file (###A###.RAW) in CTD78 format

(*.WRW,*.WSC,*.HED,*.ERR) in ASCII format

CTD 10

AT SEA COMMON USED FOR DATA ACQUISITION   CTD#10

ID   attribute 1     attribute 2     slope        bias           sens lag

PR   0.244853E-08  -0.518678E-13   0.100212E+00  -0.134196E+00   0.000000E+00

TE   0.184585E-11   0.000000E+00   0.499957E-03  -0.402169E-02   0.250000E+00

CO  -0.650000E-05   0.150000E-07   0.100456E-02  -0.143578E-01   0.000000E+00

SW   0.280000E+01   0.300000E+04   0.100000E+01   0.000000E+00   0.000000E+00

OC  -0.360000E-01   0.115000E-03   0.148000E-02   0.000000E+00   0.000000E+00

OT   0.750000E+00   0.000000E+00   0.128000E+00   0.000000E+00   0.000000E+00

TP   0.000000E+00   0.000000E+00   0.100000E+01   0.000000E+00   0.000000E+00

RT   0.000000E+00   0.000000E+00   0.000500E+00  -2.000000E+00   0.000000E+00

QU   0.000000E+00   0.000000E+00   0.100000E+01   0.000000E+00   0.000000E+00

*above are labcals, new co bias and slope input when available*

CTD 8

AT SEA COMMON USED FOR DATA ACQUISITION   CTD#8

ID   attribute 1     attribute 2     slope        bias           sens lag

PR   0.216893E-07  -0.194672E-12   0.992262E-01   0.117154E+01   0.000000E+00

TE   0.205519E-11   0.000000E+00   0.499853E-03   0.774354E-02   0.250000E+00

CO  -0.650000E-05   0.150000E-07   0.100229E-02  -0.247875E-01   0.000000E+00

SW   0.280000E+01   0.300000E+04   0.100000E+01   0.000000E+00   0.000000E+00

OC  -0.360000E-01   0.115000E-03   0.148000E-02   0.000000E+00   0.000000E+00

OT   0.750000E+00   0.000000E+00   0.128000E+00   0.000000E+00   0.000000E+00

QU   0.000000E+00   0.000000E+00   0.100000E+01   0.000000E+00   0.000000E+00

COMMENTS regarding data acquisition:

LEG 1:  Test station 999 extremely noisy, vax crashed twice due

to integer*2 problem i.e., num_errs exceeded 32,749.  Retermination

didn't fix problem.  Finally discovered underwater cable

and wiring in pigtail needed to be reconstructed.

Station 1 collected with arrays in ctd78_common.inc dimensioned

to integer*4.  This proved to be unreadable and incompatible

with the post processing ctd78 format routines.  Data was clean

however.  Aqui89 was reverted back to the original code (integer*2)

but compiled so that errors would not increment if more than

32,749.

Plot routine seems to have a bug in its initial startup.

Downcast is offset by almost 1/2 inch compared to the upcast.

During steam between 1st and 2nd group of stations (10 and 11),

lost data connection.  Underwater cable reconstructed again,

along with mecca and slippering repairs.  CTD conductivity

shift fresh (due to repairs or steam?).

Station 17: logger crashed on upcast.  Reason unknown.

LEG 2:

Station 34: logger crashed at start of cast:  this time we

ascertained that it was the defragmentation program running

as a background process that caused the crash.  The defragmenter

was removed from the VAX used for acquisition and the cast

proceeded normally.

Oxygen trace more noisy than usual.

Theta-S plot reveals conductivity noisy beneath 500m and

shifted ~.008 fresher.... Particulate matter caught in conductivity

cell?? Conductivity back to normal on subsequent casts.

Station 35: Oxygen profile very noisy.  Replaced with new unit

after this cast.

Station 36:  New Ox sensor; looks good.

BOTTOM CONTACT:

PROBLEM STATIONS: Station 1

*************************************************************************

*                                                                       *

*               S H I P B O A R D   P R O C E S S I N G                 *

*                                                                       *

*************************************************************************

DESCRIPTION OF COMPUTER SYSTEM USED:

DRAGON

VaxStation 3600

16 mb of memory

1 gigabyte disk space

IGUANA

used as backup and for post-processing

identical system to DRAGON

CTLOG

used on this cruise for copying VAX disk editted CTD78

format data files  [ctd.EN223D040.ctded78]####D###.EDT

to 9T tapes LS01-

CTDED78

run on 9T MicroVAX Acquisition CTD78 format raw data tapes.

output to MicroVAX disk files [CTD.EN223D040.CTDED78]####D###.EDT

error identification downtrace

magnetic tape serial numbers:

LS01   stations    2 -  9

LS02   stations   10 - 18

Water sample Programs:  WS_EN223

Water sample filenames:   Salinity and Oxygen Rosette Samples

En223c10s.mlt - all stations

En223c10s.DYN

Nutrient water sample filenames:

Sta#.HYD

[COMMENTS]

*************************************************************************

*                                                                       *

*                   C A L I B R A T I O N S                             *

*                                                                       *

*************************************************************************

CTD SENSOR CALIBRATIONS used for POST CRUISE PROCESSING

********************************************************************

******                                                       *******

******    CALIBRATION of PRESSURE and TEMPERATURE SENSORS    *******

******                                                       *******

********************************************************************

.CAL format calibration data files on PC

Temperature and Conductivity in C:\LABCAL\LABFIT

Pressure cals (with new tit std) in C:\LABCAL\LABFIT\RUSKCAL

CTD #10   10MR91TC.CAL  PRE-cruise TE, CO

.FIT  for listing of scaling factors and residuals

10MR91PR.CAL  PRE-cruise PR

.FIT

10MY91TC.CAL  POST-cruise TE, CO

.FIT  for listing of scaling factors and residuals

10MY91PR.CAL  POST-cruise PR

.FIT

CTD #8     8MR91TC.CAL  PRE-cruise TE, CO

.FIT  for listing of scaling factors and residuals

8MR91PR.CAL  PRE-cruise PR

.FIT

NO POST-cruise cal done as instrument was never used.

PROGRAM:    RUSKCAL.EXE

--------

PRESSURE:  cubic down (0-9000 db range)

--------

BI= -(pre-lowering deck unit PR reading)

BIAS        SLOPE         ATTR 1       ATTR 2

PRE-CRUISE  CTD10  -.134196E0   .100212E+0    .244853E-8   -.518678E-13

CTD8    .117154E1   .992262E-1    .216893E-7   -.194672E-12

POST-CRUISE CTD10  -.443002E0   .100241E+0    .996094E-9   -.366174E-13

(Pre/Post/Other) calibrations were used for final data load.

PROGRAM:    LABCAL.EXE

-----------

TEMPERATURE:  quadratic fit

-----------

BIAS         SLOPE         ATTR 1

PRE-CRUISE  CTD10  -.402169E-2   .499957E-3    .184585E-11

CTD8    .774354E-2   .499853E-3    .205519E-11

POST-CRUISE CTD10  -.614282E-2   .499557E-3    .785609E-11

CTD 8,10= temperature time lag of .25 seconds

(Pre/Post/Other) calibrations were used for final data load.

PROGRAM:    RUSKCAL.EXE

---------

PRESSURE:  cubic up (9000-0 db range)

---------

BI= -(post-lowering deck unit PR reading)

BIAS        SLOPE         ATTR 1       ATTR 2

PRE-CRUISE  CTD10  -.457274E0   .100301E+0   -.258800E-8    .627137E-14

CTD8   -.187178E1   .989444E-1    .361878E-7   -.345819E-12

POST-CRUISE CTD10  -.886706E0   .100280E+0   -.116550E-8   -.102759E-13

(Pre/Post/Other) calibrations were used for scaling Rosette data

see MLTOUT_.COM for details.

PROGRAM:    LABCAL.EXE

-------------

CONDUCTIVITY:

-------------

BIAS         SLOPE

PRE-CRUISE  CTD10  -.143578E-1   .100456E-2

CTD8   -.247875E-1   .100229E-2

POST-CRUISE CTD10  -.847797E-2   .100469E-2

*********************************************************

*******                                           *******

*******  CALIBRATION OF THE CONDUCTIVITY SENSOR   *******

*******                                           *******

*********************************************************

<CTD.CTDATA.MLTRG91>MLTRG91.EXE   (12-MAR-91)

A new MLTRG91 program was created PRE-cruise to encorporate the

SAL78 subroutine which was corrected by replacing C1535 with

42.914  (formerly 42.909).  This will increase salinity by about .005.

-------------------

INITIAL DATA LOAD:

-------------------

DATE:  Used MLTRG to compute conductivity scaling factors based

on best fit of CTD data to corresponding Rosette data

CTD 10: INITIAL load   STATIONS 2 - 10

MPLRESC10        fit for CO Bias,Slope

COSLC10          apply bias,fit for slope T < 4.0

No P edit

TE lag= .25

BIAS         SLOPE

-.12148581E-1   .10039867E-3

SECOND load   STATIONS 11 - 17

MPLRES11T17      fit for CO Bias,Slope

COSL11T17        apply bias,fit for slope T < 4.0

No P edit

TE lag= .25

BIAS         SLOPE

-.13137923E-1   .10040915E-3

Comparison of deep water theta/salt plots of CTD (M version) traces

and rosette data (EN223S.DYN) yielded the following:

Supreme results-station 2-10 overlay with 11-17.

Delta conductivity vs. station plot looks great sta 1-10 fit!  Rosette

salts slightly fresher than CTD second fit- sta 11-17.

SUGGESTIONS FOR other DATA LOADS:

<CTD.CTDATA.PRSORT91>PRSORT91.EXE   (12-MAR-91)

Program PRSORT91 was used to convert time series CTD78 format data

to VAX disk files of pressure averaged observations at 2 db intervals.

The '.CTD' data files include variables OC and OT from which oxygen

calibrations will be refined later.

A new PRSORT91 program was created PRE-cruise to encorporate the

new titanium pressure sensor and new pressure temperature algorithm.

See the subroutine PRESSC.FOR.  This sensor should not exhibit any

hysteresis.  Also, the SAL78 was corrected, replacing C1535 with

42.914  (formerly 42.909).  This will increase salinity by about .005.

The following FINAL command files were used to run PRSORT loading

data from CTDED CTD78 FORMAT tapes:

PSLS01.PST  ctd #10  stations:   2-9

PSLS02.PST  ctd #10  stations:  10-18

************************************************

******                                    ******

******    CALIBRATION of OXYGEN SENSOR    ******

******                                    ******

************************************************

Basic Method for generating oxygen fits:

----------------------------------------

<CTD.CTDATA.EXE>HYDROUT    (21-may-85)

used to transform ENCALLS.MLT to ENCALLS.DYN (NODC format)

LISTGRID produces ASCII listing (STALEG#.LST) of Rosette

data from '.DYN' file

<CTD.CTDATA.EXE>MLTOXDWN   (31-mar-86)

used to generate FOR081.DAT file (saved as OXDNsta#.MLT)

of water sample data from the appropriate SUBINDEX.CTD

file and SHCRU.DYN file by extracting data scans from

CTD downtrace profiles which correspond to Rosette

sampling pressures.

<CTD.CTDATA.EXE>OXFITEDC   (1-feb-84)

run on each FOR081.DAT file to arrive at scaling

parameters for the oxygen sensor; fit the least number

of parameters necessary to reduce s.d. to +/- 2.8

<CTD.CTDATA.EXE>OXCALCL    (26-jun-86)

used to apply oxygen scaling factors derived by OXFITEDC

to PRSORTed  .CTD files.  Outputs 'X' version.

$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$

$$$$$$                                                               $$$$$$

$$$$$$          PROGRESSION OF OXYGEN SENSOR CALIBRATIONS            $$$$$$

$$$$$$                                                               $$$$$$

$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$

--------------------------

I....  DATE :   'X' versions created from *M* files

--------------------------

THESE 'X' versions were created with the following fits generated

to the station bounds listed :

MLTOXDWN used to generate a series of .MLT files for oxygen

calibration groups; Br 7  FOR081.DAT output files renamed to OXDNsta#.MLT

OXDN2.MLT       CTD #10  Oxygen sensor A

OXDN3.MLT

OXDN4.MLT

OXDN5.MLT

OXDN6.MLT

OXDN7.MLT

OXDN8.MLT

OXDN9T10.MLT

OXDN11.MLT

OXDN12.MLT

OXDN13.MLT

OXDN14.MLT

OXDN15.MLT

OXDN16.MLT

OXDN17.MLT

OXFITEDC run on each OXDN...MLT file separately to arrive at scaling

parameters for oxygen sensor as follows:

fit 4, 5 OR 6 parameters, starting with the first until sd +/- 2.8 reached

set OXMIN to .001 (default is .4 but this min value must be changed to

account for much lower values in the PACIFIC OCEAN)

OXOC2.DAT    CTD #10  Oxygen sensor A   applied to sta 2

OXOC3.DAT                               applied to sta 3

OXOC4.DAT                               applied to sta 4

OXOC5.DAT                               applied to sta 5

OXOC6.DAT                               applied to sta 6

OXOC7.DAT                               applied to sta 7

OXOC8.DAT                               applied to sta 8

OXOC9t10.DAT                            applied to sta 9-10

OXOC11.DAT                              applied to sta 11

OXOC12.DAT                              applied to sta 12

OXOC13.DAT                              applied to sta 13

OXOC14.DAT                              applied to sta 14

OXOC15.DAT                              applied to sta 15

OXOC16.DAT                              applied to sta 16

OXOC17.DAT                              applied to sta 17

OXOC18t22.dat                           applied to sta 18-22

OXOC23.DAT                              applied to sta 23

OXOC24.DAT                              applied to sta 24

OXOC25.DAT                              applied to sta 25

OXOC26.DAT                              applied to sta 26

OXOC27.DAT                              applied to sta 27

MIN/MAX STA,#    2.    2.  1 STD DEV.=0.10562E+00 #OB   22  DOX=0.296

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.135 1.101  0.1277E-03 -0.4162E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#    3.    3.  1 STD DEV.=0.13965E+00 #OB   21  DOX=0.391

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.382 1.325  0.1895E-03 -0.4576E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#    4.    4.  1 STD DEV.=0.71876E-01 #OB   23  DOX=0.201

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=   0.023 0.961  0.9973E-04 -0.4202E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#    5.    5.  1 STD DEV.=0.64862E-01 #OB   23  DOX=0.182

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.059 0.987  0.1265E-03 -0.2898E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#    6.    6.  1 STD DEV.=0.90862E-01 #OB   22  DOX=0.254

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=   0.638 0.555 -0.3662E-05 -0.8239E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#    7.    7.  1 STD DEV.=0.41485E-01 #OB   19  DOX=0.116

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=   0.135 0.879  0.6720E-04 -0.4563E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#    8.    8.  1 STD DEV.=0.66391E-01 #OB   23  DOX=0.186

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.172 0.872  0.1989E-03  0.2899E-01  0.4372E+00  0.7918E+01

MIN/MAX STA,#    9.   10.  2 STD DEV.=0.62033E-01 #OB   37  DOX=0.174

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=   0.147 0.797  0.1104E-03 -0.4107E-01  0.6307E+00  0.8000E+01

MIN/MAX STA,#   11.   11.  1 STD DEV.=0.61136E-01 #OB   23  DOX=0.171

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.123 1.007  0.1302E-03 -0.3598E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#   12.   12.  1 STD DEV.=0.66469E-01 #OB   23  DOX=0.186

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.093 0.974  0.1255E-03 -0.3236E-01  0.5204E+00  0.7996E+01

MIN/MAX STA,#   13.   13.  1 STD DEV.=0.25592E-01 #OB   23  DOX=0.072

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.056 0.951  0.1215E-03 -0.3519E-01  0.4172E+00  0.8004E+01

MIN/MAX STA,#   14.   14.  1 STD DEV.=0.70339E-01 #OB   22  DOX=0.197

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.167 1.052  0.1503E-03 -0.3577E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#   15.   15.  1 STD DEV.=0.55334E-01 #OB   23  DOX=0.155

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.031 0.915  0.1195E-03 -0.3270E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#   16.   16.  1 STD DEV.=0.94396E-01 #OB   23  DOX=0.264

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=   0.046 0.932  0.9452E-04 -0.4620E-01  0.7808E+00  0.8021E+01

MIN/MAX STA,#   17.   17.  1 STD DEV.=0.11408E+00 #OB   24  DOX=0.319

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.114 0.890  0.1538E-03  0.3355E-02  0.7500E+00  0.8000E+01

MIN/MAX STA,#   18.   22.  5 STD DEV.=0.56451E-01 #OB   93  DOX=0.158

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=   0.022 0.909  0.1070E-03 -0.3120E-01  0.1848E+00 -0.1028E+01

MIN/MAX STA,#   23.   23.  1 STD DEV.=0.57601E-01 #OB   14  DOX=0.161

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.261 1.072  0.1821E-03 -0.4640E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#   24.   24.  1 STD DEV.=0.10091E+00 #OB   14  DOX=0.283

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.380 1.184  0.1570E-03 -0.3119E-01  0.7500E+00  0.8000E+01

MIN/MAX STA,#   25.   25.  1 STD DEV.=0.31720E-05 #OB    4  DOX=0.000

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=   0.080 0.826  0.2663E-03 -0.4737E-01  0.1831E-01  0.7996E+01

MIN/MAX STA,#   26.   26.  1 STD DEV.=0.17535E-01 #OB    6  DOX=0.049

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=   0.301 0.673  0.5207E-04 -0.6368E-01  0.9054E-01  0.8000E+01

MIN/MAX STA,#   27.   27.  1 STD DEV.=0.57721E-01 #OB    7  DOX=0.162

OX BAIS OC SL    PCOR         TCOR      WT  OXLAG

OX PARMS=  -0.363 1.152  0.5528E-03 -0.1256E-01 -0.8698E+00  0.8000E+01

Comparison of full profile pressure/O2 plots of CTD traces

and rosette data (  .DYN) yielded the following:

Each station fit to itself initially, trouble fitting in deep

water.  Single fits still not ideal.

**************************************************

*********                                *********

*********        QUALITY CONTROL         *********

*********                                *********

**************************************************

<CTD.CTDATA.EXE>CTDED.EXE  (27-oct-86)           EN223CTDED.COM

To correct erroneous data scans (i.e. interpolate over pressure

bounds to get rid of salinity and oxygen spikes in the data)

Outputs 'E' version.

STATION       VARIABLE       ACTION        P1, P2      COMMENTS

------------     --------    -----------     ---------    --------------------

2             SA           19/         3625,3633    INTERPOLATE

_______________________________________________________________________________

<CTD.CTDATA.EXE>CRSUMARY1.EXE   (12-jan-83)

To create SUBINDEX.CTD file of selected versions of CTD data files

FINAL ARCHIVED VERSIONS OF SUBINDEX.CTD

DATES   1991              SUBINDEX.CTD   final XE version data files

1991              MSUBINDEX.CTD  prsort 'M' versions

1991              FSUBINDEX.CTD  final XE versions

(copy of subindex.ctd)

...............................................................................

<CTD.CTDATA.EXE>INDEXED.EXE   (21-aug-84)

To correct erroneous header information in SUBINDEX.CTD files

and transfer that corrected info to the station data files.

branch 3:  auto-check mode to scan for errors, run interactively

branch 1:  to correct header info in SUBINDEX.CTD file

branch 2:  transfer header info

from SUBINDEX.CTD file to data files

Header information corrections:

STATION       CATEGORY         CORRECTION         COMMENTS

------------     --------    -------------------     -------------------------

All stations were corrected with ADCP (Acoustic Doppler Current Profiler) data

[ctduser.jmt1.nerctape]readpostion.exe

reads in NAV file from 9T

outputs POSITION#.dat file

where.exe

reads in postion#.dat file

outputs whereout#.dat file

(which can be included in MW893INDEXED.com

DATE         COMMAND FILE      DOCUMENTATION       COMMENTS

-----------      ------------      -------------       -------------------------

INDEXED.COM         EN223INDEXED.LST   all stations, first pass,

branch 1 only

INDEXTRNSFER.COM    INDEXTRNSFER.LST     all stations, branch 2

INDEXEDB3.DAT        all stations, branch 3

**************************************************

*********                                *********

*********           B A C K U P          *********

*********                                *********

**************************************************

Command file:  @CTDA:<CTDSOFT.COMS>BACKUP    BLOCK=10000  BPI=1600

Run interactively.

DATE         TAPE LABEL         WHERE STORED      SAVE_SET NAME(S)

------       ------------       --------------     ----------------

DA-MON-YEAR       SHCRU            CLARK/MCLEAN      1. SHCRUPROJ.FIN

2. SHCRUPROJ.BCK

**************************************************

*********                                *********

*********             N O D C            *********

*********                                *********

**************************************************

<CTD.CTDATA.EXE>NODCEXCH.EXE  (29-may-85)

WHERE       QUALITY CHECK

DATE         ALL STAS?    DENSITY      STORED      G HEIMERDINGER

----         ---------    -------      ------      --------------

********************************************************

*********                                      *********

*********        P U B L I C A T I O N S       *********

*********                                      *********

********************************************************