OCEANUS CRUISE 218

HRP PROFILER / CTD  CRUISE

6 DECEMBER 1990 :

CTD #9 WAS USED FOR THIS CRUISE.

PRECRUISE CALIBRATIONS WERE COMPLETED AND THEY LOOKED FINE.

IT WAS DISCOVERED (AFTER THE CRUISE) THAT THE CONDUCTIVITY CELL WAS CRACKED

AND HAD BEEN AT ONE TIME REPAIRED (THIS HAPPENED PRIOR TO THE DAYS OF GARY

BOND).

PRECRUISE CALS DESCRIBED THE THIS SENSOR PERFECTLY FOR THE FIRST TWO CASTS.

LAB CALS DON'T SUBJECT THE SENSOR TO PRESSURE HOWEVER.  THE CRACKED CELL

PRESUMEABLY FALTERRED UNDER THE PRESSURE STRAINS OF THE FIRST 2 CASTS.

THERE WAS NO POST CRUISE CALIBRATION ON THIS INSTRUMENT.  AT THE TIME,

ALL INSTRUMENTS WERE BEING REFURBISHED BY THE NEW CTD GROUP, AND UPON

DISCOVERY OF THE CRACKED CONDUCTIVITY CELL, IT WAS REPLACED WITH A HEALTHY

ONE.

THE VAX CTDED EDITTOR WAS USED TO CORRECT DATA:

STATIONS 1 AND 2 WERE RIGHT ON WITH THE HISTORIC (OC134, AT109) THETA/S

CURVES AT SEA USING THE PRE-CRUISE CALIBRATION INFORMATION.  A HUGE SHIFT

SALTY (0.055) OCCURRED DURING STATIONS 3 AND 4 AND VARIED SLIGHTLY FROM

THEN ON THROUGHOUT THE CRUISE.

THE VAX CTDED EDITTOR WAS USED TO SUBTRACT A BIAS (TOP TO BOTTOM) FROM THESE

DATA.  IT WAS DECIDED THAT A BIAS ADJUST (AS OPPOSED TO A SLOPE ADJUST)

SHOULD BE DONE FOR SUCH A LARGE CONDUCTIVITY SHIFT (OTHERWISE, THE OFFSET

IN THE SURFACE WATER WOULD BE ASTRONOMICAL).