CTDA:[CTD.TW016P003]P16STATUS.DOC

STATUS OF P16C

as of 8/92.  CAM.

WOCE WATERSAMPLE FILE

In April, bottle data was reprocessed with post-cruise pressure

and temperature cals and a post cruise revision of the .sox files.

In August, 9 stations were reordered/modified and processed again.

The latest version of this file resides in CTDA:[CTD.TW016P003.HYDRO]

as TW16FNL.SEA.  New hydro .sox files for these 9 stations still need

to be merged in, along with any other fixes to quality flags as noted

in Lynne's mail message.

The changes made for the 9 stations are as follows:

STA 252- reorder bottles 1-24,36,25-35

STA 259- cast 5

STA 267- bottles 1-9 changed to 1,2,4-10

STA 274- reorder bottles 1-23,36,25-35

STA 278- reorder bottles 2-24,1,25-36

STA 280- reorder bottles 2-24,1,25-36

STA 283- delete tags 4 & 9, reorder bottles 1-9

STA 286- reorder bottles 1-24,36,25-35

STA 289- bottle 29 changed to bottle 19

These changes were made to the .wrw raw aqui tag files,

which were transformed into .bt2 files with BTLFMTVX, and

then into .tmp files using WOCTMPV2, and finally into the .sea

file using SEAMERG2.  Please note, that these programs reflect

the latest woce format.  Btlfmtvx, woctmpv2, seamerg2 were

created post-cruise and are different from the at-sea programs.

These programs reside in ctda:[ctdev.ws...].  The .wrw,.bt2, and

.tmp files reside in ctda:[ctd.tw016p003.btldata].  The tmp files

were also transferred over into ctda:[ctd.tw016p003.hydro] which

were then merged with the .sox files into the .sea file.

The TW16FNL.SEA file does not however have the final pressure,

ctd salinity or ctd oxygen in it.  On leg 1 of P6, Bob Millard determined

that some of the pressure temperature cals were in error.  It was

also determined that the aqui89 program was often writing erroneous

PT numbers in the .wrw file.  John Toole has written a program to go

back into the .RAW (raw data files from aqui89) and extract the proper

value for pressure temperature at each tag.  This means that the

.RAW files need to be reloaded on to disk.  Once new .wrw files

are created (with the correct PT values), the data needs to be reprocessed

with the new PT cals for each instrument.  Jane Dunworth is working on

a new SEAMERGE program so that the P column can be overwritten.

Once the final cals for conductivity and oxygen are determined,

the ctd salinity and oxygen columns need to be overwritten.

CTD DATA

The ctd data needs quite a bit of scrutiny before it can be post-

processed.  P16C.DOC will be a useful starting point to see what

the data was processed with at sea.  Upon post-cruise analysis, it

was discovered that some of the stations (both with ctd 9 and 10)

revealed what we are calling "conductivity hysteresis."  These

stations showing this characteristic were re-pressure-sorted using

the final post-cruise pr and te cals and nominal co cals in

ctda:[ctd.tw016p003.updown]*.ctd.  Please note that this may need

to be redone with the correct PT cals.  This should be relatively easy

and quick by simply changing the PSTW*.PST command files.

The program UPDWNCOND.for was modeled after MW893 updowntheta.for

to reveal the up/down conductivity difference.  It is still in its

development stage and needs to be debugged.  The idea is to create

a new .wsd co fitting file with downtrace conductivity in it so

that the ctd downtrace data for those stations showing hysteresis

can be fit to it.  So, for the stations showing hysteresis, two sets of

conductivity cals will be used, one set to process the uptrace watersample

.sea file, and a different set to prsort the downtrace .ctd data files with.

GOOD LUCK.