Two NBIS CTD/O2 instruments were employed on this cruise.   WHOI CTD

#10 (serial # ______) served as the primary instrument and CTD # 9 (serial #

______) as the backup instrument. Both were equipped with sensors to measure

pressure, sea water temperature, redundant temperature, conductivity, and

dissolved oxygen concentration.  The accuracy of the reduced

data set hinges on the calibrations of these sensors.  Both laboratory

measurements and water sample data obtained at sea were used to calibrate the

sensors.  General information on CTD calibration and data processing

procedures can be found in the reports of Fofonoff, Hayes and Millard (1974)

and Millard and Galbraith (1982).

Laboratory calibrations, performed before and after the P16C cruise,

provide the sole correction information for the CTD pressure and temperature

sensors.  Note that the temperature and pressure calibrations are used to

scale the data profiles as well as the CTD component of the rosette water

sample data files.

(A note on the accuracy of the laboratory calibrations!!)

No electronic adjustments were made to the temperature sensor interface

boards during laboratory calibrations in order to preserve a long standing

history on the stability of these sensors.  Instead, corrections, determined

by polynomial least-squares fits to laboratory calibration data, were applied

to the data.  Temperature calibrations consisted of quatdratec fits to 8-24

temperature points ranging between 0 and 30 degrees C in reference to a

platinum thermometer standard (Figure *** showing pre and post TE cal plots

for each CTD).  The following temperature correction algorithms were used in

the reduction of CTD downcast and water sample rosette data collected on this

cruise.

(List temp algorithms for CTD's 9 and 10)

(Note here, too, the time lags used)

Titanium pressure sensors were used on both CTD instruments for this

cruise, thus significantly reducing the down/up pressure historesis commonly

found in the past.  Pre and post cruise pressure calibrations were

completed using a

dead weight tester in the laboratory;  data were sampled at 1000 psi

intervals with both increasing and decreasing pressure between 0 and 10000 psi.

Data reduction employed a quadratic calibration algorithm determined from a

least squared fit to these data (see laboratory PR calibration plots for

CTDs 9 and 10).  For titanium pressure sensors, the laboratory calibration

pressure bias terms are thought to be more reliable than pressure bias terms

determined from the  pre-lowering deck unit pressure reading.  The following

quadratic down-up (o-9000db) pressure calibration algorithm was applied to

CTD #10 data:

The quadratic down-up (0-9000db) calibration algorithm for instrument #9

pressure data was:

(When talking about CO cals be sure to mention abnormal Beta terms

used for all 3 sensors employed on this cruise!!!)