## maintenance licor SiSt
date | task
2017-09-28 | Changed fluke coëfficients in DAS
| OLD | <sensor x0="-201,6801" x1="1,520803" x2="0,004878748">fluke</sensor>
| NEW | <sensor x0="-233,0151" x1="2,095256" x2="2,250497E-03">fluke</sensor>
2017-10-13 | replaced chemicals of scrubber
2017-11-08 | replaced 800ppm CO2 gas => 796,34
2017-11-22 | licor flooded
2017-11-23 | clean/dry licor + dry pump + connect pump equilibrator correctly (tubing of pump was 'geknakt')
27-11-2017 tot 4-12-2017 test script
2018-01-08 | install new licor
2018-01-16 | install sollinoid (A4) vent for calibration gasses, so they do not need to go trough the equilibrator
wires of sollinoid A1 were loose in electrical cabinet => terug vastgemaakt
2018-01-17 | take sollinoids (A2 and A3) before and after licor out of script => board of Michiel crashed, voeding (6 Ampère) is niet sterk genoeg om 4 sollinoids samen te laten werken
2018-01-18 | Sollinoids A2 and A3 completely removed.
2018-01-30 | licor opened and checked if there was water/salt inside. Afterwards closed everything again and placed everything back.
flow meter before equilibrator did not work anymore
2018-02-01 | replaced 400ppm CO2 gas => 400,29
cleaned flow meater before equilibrater => it works again
changed pressure of calibration gas that goes through the licor from 0,5 to 1,0 (flow of water that goes through equilibrator is between 1,0 and 1,2 same as gasflow that is measured after it comes out of the licor (1,2))
2018-02-07 | flow meter was blocked again.
2018-02-09 | changed the flow path of calibration gasses so they do not go through the nafion tube anymore.
Installed a sollinoid (normal open) after the nafion tube that is open when normal gasses pass and that is closed when calibration gasses pass.
adjusted flow of calibration gasses again to 1, after changeing the system the flow of the gasses was at 4 (nafion tube blocked???)
2018-03-02 | Removed fluke+pt100 for calibration, started calibration at around 16:00
2018-03-05 | Reinstalled fluke+pt100, calibration finished except for 25°C setpoint
2018-03-20 | changed flow of calibration gasses from 1 to 0,1 => did check => stdev were not ok
2018-03-21 | changed flow of calibration gasses from 0,1 to 0,3 => did check => left flow of calibration gasses at 0,3
2018-03-25 | new coefficients of fluke + pt100 inserted in DAS
2018-04-03 | changed script to do check every 3 hours
2018-04-05 | changed equilibrator in ship and inserted pt100 in the equilibrator
2018-05-14 | changed equilibrator in ship
2018-05-30 | Stopped sytem and got everything out the ship.
2018-06-25 | system is +- operational again on the schip
changed 0 gas bottle to a new one
Coëfficients pt100
| pt100 equilibrator |  | <sensor x2="4,90E-14" x1="0,0000158" x0="-244,0587953">marbles</sensor>
| pt100 drain |  | <sensor x2="5,13E-14" x1="0,0000157" x0="-243,2498364">marblesdrain</sensor>
2018-07-03 | Switched Licors
2018-07-25 | changed 400ppm CO2 gas bottle, was empty => new value 394,1
2018-07-26 | regulated 400 gas flow to 0,5
2018-08-23 | 250 gas bottle empty,
2018-08-27 | 250 changed gas bottle=> new value 249.0
2018-09-03 | 800 gas bottle empty
2018-09-04 | 800 changed gas bottle => new value 792.3
2018-09-27 | 0 gas bottle empty
2018-09-28 | 0 gas bottle changed
regulated flow calibration gasses + checks to +-0.2
2018-10-23 | switched twoo regulators from the calibration gasses with 2 new ones (0 and 800 gas)
2018-11-09 | changed timing of the checks in the script of DAS
2018-11-01 | 0 bottle is empty
2018-11-13 | placed 0 gas bottle (air products) inside the dry lab next to the system
2018-11-14 | replaced the black tubing that goes to the vici with BEV A line tubing (some water drops fell out of the black tubing)
can explain why we have such strange values suddenly.
2018-11-29 | changed regulator on 0 gas bottle inside the ship
2018-12-10 | closed the door of the dry lab

## cleaning-cal optode ship
date | task
2018-02-01 | optode 2194 sooguard cleaned
2018-04-05 | optode 2194 removed from sooguard and sooguard was rinsed
2018-04-16 | optode 2194 cal and reposition in bow thruster room
optode (1245 calibration on CTD)
2018-07-23 | optode 2194 sooguard cleaned
2018-09-21 | optode 2194 sooguard cleaned
2018-11-16 | optode 2194 sooguard cleaned
2019-01-08 | optode 2194 removed

## SBE21
date | task
2017-09-27 | placed SBE 21 S/N:3257 in drylab
2018-02-01 | SBE 21 cleaned S/N: 3257
2018-04-05 | SBE 21 cleaned S/N: 3257
2018-07-23 | SBE 21 cleaned S/N: 3257
2018-09-21 | SBE21 S/N: 3257 changed by SBE21 S/N: 3385
2018-11-16 | SBE21 cleaned S/N: 3385

## SBE38
date | task
2017-02-10 | SBE38 S/N: 3840702-0267 placed in bow thruster room
2018-09-28 | SBE38 S/N : 3840702-0267 replaced by SBE38 S/N: 38-0995